## Cover sheet
Maintenance handbook for BSL-2 and BSL-3 laboratories
In accordance with ContainO and PEMO
Appendix 1
Safety level 3 laboratories
Excel tables are provided as an Appendix to the maintenance handbook. They show the requirements for the maintenance and inspection journals, and give explanations and examples for consistent further development.
Maintenance and inspection journals are given for all safety-relevant systems, structured according to safety level 2 and 3 laboratories. They may be completed with reference to the SWKI guidelines or in collaboration with the specialist companies.
The Excel tables are weighted towards ventilation systems, as experience has shown that these are the priority for maintenance and inspection journals.

## Journal example
| Maintenance and inspection |  | Name of system / component
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning
Function |  | Maintenance / inspection plan
|  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks
Reliability
Safety measure
Disposal
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Schedule example
|  | Maintenance and Inspection Schedule | [46 empty] | Year
|  | Company: | [20 empty] | Location:
|  | Person responsible: | [20 empty] | Facility:
|  | Place / date: | [20 empty] | Commissioning
Code | Part | Week | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52
carried out by / signed: ……………………………..
Date: ……………………………..

## Surrounding surfaces
| Maintenance and inspection |  | Surfaces surrounding BSL-3 lab areas
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Integrity is necessary for the surfaces surrounding BSL-3 laboratories, so that they may be fumigated / decontaminated. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Walls | [12 empty] | more detail needed
Check surface conditions of epoxy resin on concrete wall / floor for damage Check door seal (inflatable or rigid) for damage Check windows (e.g. bulletproof, fire protection class etc.) for damage Check seals to wall penetrations for damage |  |  | Ceiling | [12 empty] | in work sheets for
|  |  | Floor | [12 empty] | individual parts
Doors
Windows
Reliability |  |  | Penetrations
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
Observe manufacturer's instructions.
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Wastewater deactivation
| Maintenance and inspection |  | Wastewater deactivation plant
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
BSL-3 laboratories should generally be run as dry labs and not as wet labs. If a wastewater deactivation plant is necessary, it may be desgined as a thermal or chemical installation, depending primarily from the use of the laboratory for microbiology. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Wastewater deactivation | [12 empty] | more detail needed
Check storage tank, catch basin and pipework for damage, corrosion and fastening Check for watertightness Check insulation for damage Check drains for corrosion and damage Check temperature and pressure gauges for damage and accuracy of display Check overpressure relief valve with HEPA filter for functioning and watertightness |  |  | Effluent tanks | [12 empty] | in work sheets for
|  |  | Water pipes and accessories | [12 empty] | individual parts
Shut-off device
Anti-flooding devices
Reliability |  |  | Drainage pump
Safety can be ensured only if the supplier's instructions are observed. |  |  | Catch basin
Decontamination
Valve controls
HEPA filter
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Service water
| Maintenance and inspection |  | Service water network
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A holding tank should be provided for mains separation between drinking water and service water (BSL-3 laboratory separate from remaining drinking water supply), according to the state of technology. A pressure booster and water heating for service water for lab sinks, eye showers and emergency showers may be needed, depending on use. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Holding tank | [12 empty] | more detail needed
Check holding tank, water pipes and accessories for damage, (external) corrosion and fastening Check insulation for damage and completeness Check for watertightness (visual check) Check sections of pipes for incrustation Check pressure booster, pressure vessels for damage, (external) corrosion and fastening Check pressure vessels for watertightness Check functioning of shut-off device and safety equipment Check pressure relief and shut-off valves in the expansion line (open position, securing) |  |  | Water heater | [12 empty] | in work sheets for
|  |  | Pressure booster | [12 empty] | individual parts
Water pipes and accessories
Shut-off device
Reliability |  |  | Service water pump
Safety can be ensured only if the supplier's instructions are observed. |  |  | Eye showers
Emergency showers
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Compressed air
| Maintenance and inspection |  | Compressed air supply
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A compressed air supply system generally consists of compressors, buffer tanks, pressure monitoring equipment, compressed gas cylinders, automatic switching station, removal station, and where applicable, HEPA filter |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Compressors | [12 empty] | more detail needed
Check for damage Check oil level Top up oil Replace oil Check functioning of compressor Check functioning of switches, safety devices and reducing station Adjust switches, safety devices and reducing station Check functioning of automatic drainage Drain boiler and reducing station Check filter for dirt Clean filter Replace filter Check functioning of dehumidifier Clean |  |  | Buffer tanks | [12 empty] | in work sheets for
|  |  | Pressure monitoring | [12 empty] | individual parts
Compressed gas cylinders
|  |  | Automatic swtiching stations | [12 empty] | In accordance with
Reliability |  |  | Removal station | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  |  | Where applicable, HEPA filter | [12 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## CO2 supply
| Maintenance and inspection |  | Carbon dioxide supply
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Carbon dioxide supply systems generally consists of compressed gas cylinders, automatic switching station, gas failsafe device, removal station, non-return valve, and, where applicable, HEPA filter |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Compressed gas cylinders | [12 empty] | more detail needed
Check installation for damage, dirt, corrosion Check oil level Top up oil Replace oil Check functioning of compressed gas cylinders Check functioning of gas failsafe device Check functioning of switches, safety devices and reducing station, and adjust where applicable Check filter for dirt Clean filter Replace filter |  |  | Automatic switching station | [12 empty] | in work sheets for
|  |  | Gas failsafe device | [12 empty] | individual parts
Removal station
|  |  | Non-return valve | [12 empty] | In accordance with
Reliability |  |  | Where applicable, HEPA filter | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Liquid nitrogen
| Maintenance and inspection |  | Liquid nitrogen supply
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Liquid nitrogen supply systems generally consist of the following components: Tank installation with filling level and pressure monitoring, pipes, shut-off valve and an oxygen deficiency warning system in the lab area. The filling station is situated outside containment |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Tank installation | [12 empty] | more detail needed
Check tank installation for damage, dirt, corrosion Check functioning of filling level monitoring Check functioning of pressure monitoring Check pipes and check functioning of insulation Check functioning of oxygen deficiency warning system |  |  | Filling level monitoring | [12 empty] | in work sheets for
|  |  | Pipes | [12 empty] | individual parts
Oxygen deficiency warning system
|  |  | Filling station | [12 empty] | In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Fire extinguishing
| Maintenance and inspection |  | Fire extinguishing system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
National legislation on fire safety requires fire extinguishers to be provided. Sprinkler systems should be avoided if possible because of contamination through extinguishing water. In some cases, coordination with the fire service responsible is necessary. Should the fire service require sprinkler or spraymist systems in addition to fire extinguishers, a wastewater deactivation plant should be planned for the extinguishing water. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Fire extinguishers | [12 empty] | In accordance with
Check sprinkler for coatings Check sprinkler settings Check pressure vessels for tightness Check functioning of shut-off device and safety equipment Check pressure relief and shut-off valves in the expansion line Check pressure pads Check manometer for damage, check functioning of display Check functioning of safety valve Check feed pump for damage, corrosion, fastening, noise and tightness |  |  | Sprinkler systems | [12 empty] | manufacturer's instructions
Spraymist systems
Reliability
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Steam generator
| Maintenance and inspection |  | Steam generator
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A separate steam generator may be needed for autoclaves, if they do not have their own electric steam generation. The steam quality must be coordinated with water treatment. Heat extraction between untreated and pure steam generation may be used for autoclaves to save energy. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Untreated steam generator | [12 empty] | more detail needed
Check steam generator for dirt, damage and corrosion Check functioning of solenoid valve Check functioning of heating rods Check functioning of water supply, and check water level Check functioning of control valve Adjust stuffing box of control valve Check functioning of steam distributor and condensation line |  |  | Pure steam generator | [12 empty] | in work sheets for
|  |  | Heat exchanger | [12 empty] | individual parts
Stainless steel pipes
|  |  | Condensation line | [12 empty] | In accordance with
Reliability |  |  | Burner | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  |  | Burner fan | [12 empty] | In accordance with Swiss
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## AC supply air
| Maintenance and inspection |  | Air-conditioning system (AC) Supply air
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Conditioning of external air, temperature regulation of supply air to room conditions, regulating humidity of supply air to room conditions, filtering supply air Supply air systems possibly with redundancy (if user requires); air-conditioning systems (heating, cooling, humidifying, dehumidifying); heat recovery; silencing; supply air filter (2-level); pressure monitoring for filters and ducts; cooling pipes; volume flow controller for supply air. Users of BSL-3 laboratories may require a redundant second supply air system. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  | L01 | Device / monobloc housing | [12 empty] | more detail needed
Check underpressure in lab area Check condensate tray and droplet separator for dirt, corrosion and functioning Check directed air current into laboratory Check fan for dirt, damage and corrosion Clean fan parts and drains that are in contact with air to preserve function Lubricate bearings Check air cooler air/liquid for dirt, damage and corrosion on the air side Check functioning of flow and return Check functioning of water drain and trap |  | L11 | Coarse dust filter | [12 empty] | in work sheets for
|  | L12 | Fine dust filter | [12 empty] | individual parts
|  | L21 | PWW fan heater
|  | L24 | Air cooler | [12 empty] | In accordance with
Reliability |  | L33 | Heat recovery grid system | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  | L42 | Radial fan | [12 empty] | In accordance with Swiss
|  | L44 | Motor | [12 empty] | Standard
|  | L45 | Belt drive
|  | L46 | Fan/motor bearings
Safety measure |  | L47 | Frequency converter
|  | L57 | Steam humidifier
|  | L74 | Volume flow controller
|  | L81 | Air duct system
|  | L82 | Air intake opening
Disposal |  | L83 | Weatherproofing
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal. |  | L84 | Air outlet
|  | L88 | Fire damper
|  | L92 | Electrical controls
|  | L93 | Measurement/control peripherals
Notes |  | L94 | Switchgear combination
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## AC exhaust air
| Maintenance and inspection |  | Air-conditioning system (AC) Exhaust air
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Exhaust air systems for BSL-3 laboratories should be designed with redundancy. The exhaust air system should include heat recovery; silencing; exhaust air filter to protect heat recovery; pressure monitoring for filters and ducts; cooling pipes; volume flow controller for exhaust air |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  | L01 | Device / monobloc housing | [12 empty] | more detail needed
Check heat exchanger for dirt, damage and corrosion Check air filters for inadmissible dirt and damage (leakage); replace the air filter in question if filter stage was last replaced less than 6 months ago; otherwise replace the entire filter stage Check differential pressure; if necessary replace filter stage Check volume flow and mixer control for dirt, damage and corrosion Check functioning of volume flow controller |  | L11 | Coarse dust filter | [12 empty] | in work sheets for
|  | L33 | Heat recovery grid system | [12 empty] | individual parts
|  | L42 | Radial fan
|  | L44 | Motor | [12 empty] | In accordance with
Reliability |  | L45 | Belt drive | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  | L46 | Fan/motor bearings | [12 empty] | In accordance with Swiss
|  | L47 | Frequency converter | [12 empty] | Standard
|  | L74 | Volume flow controller
|  | L81 | Air duct system
Safety measure |  | L83 | Weatherproofing
|  | L84 | Air outlet
|  | L88 | Fire damper
|  | L92 | Electrical controls
|  | L93 | Measurement/control peripherals
Disposal |  | L94 | Switchgear combination
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## HEPA Filter for exhaust air
| Maintenance and inspection |  | AC system: HEPA filter ffor exhaust air system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
BSL-3 laboratories should have at least a separate particulate filterbox (HEPA) for the exhaust air side.HEPA filter class may be H13 or H14 or higher, depending on work done in lab.Prefilters should be provided to protect the HEPA filter and extend its operating life. There should be pressure monitoring for the filters.Gastight or airtight flaps with fumigation ports or filter replacements / filter testing devices should be provided. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  | L01 | Device / monobloc housing | [12 empty] | more detail needed
Check prefilter and HEPA filter for inadmissible dirt and damage (leakage) Replace the air filter in question if permissible differential pressure has been reached, or if filter stage was last replaced less than 6 months ago. Then replace the entire filter stage Before replacing filter, decontaminate by fumigation or replace filter through BIBO (“Bag-in Bag-out”) Check filter housing for dirt, damage and corrosion Check differential pressure in prefilter and main filter (HEPA) Check gastight or airtight shut-off valves for damage and tightness Check biohazard warning sign |  | L11 | Coarse dust filter | [12 empty] | in work sheets for
|  | L12 | Fine dust filter | [12 empty] | individual parts
|  | L13 | HEPA filter
|  |  | Gastight flaps | [12 empty] | In accordance with
Reliability |  |  | BIBO filter exchange | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  |  | Filter scanning device | [12 empty] | In accordance with Swiss
|  | L81 | Air duct system | [12 empty] | Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Refrigeration
| Maintenance and inspection |  | Refrigeration for air coolers in AC system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Cold water generation with a recooling plant (and possibly split units) should be provided for temperature control in the laboratory; possibly also with redundancy if user requires. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  | K01 | Alarm system | [12 empty] | more detail needed
Check piston, screw and turbo compressors for external dirt, damage and corrosion Check fastening and running noise Measure suction pressure Measure suction gas temperature before compressor Measure final compression pressure Measure final compression temperature at the pressure port Check oil level at the sight glass Check acidity of oil (acid test) Change oil Measure oil pressure Adjust oil pressure Measure oil temperature before and after oil cooler Measure water temperature before and after oil cooler Check functioning of oil trap |  | K02 | Energy meter | [12 empty] | in work sheets for
|  | K12 | Piping / fittings | [12 empty] | individual parts
|  | K21 | Expansion devices
|  | K24 | Pressure expansion tank | [12 empty] | In accordance with
Reliability |  | K41 | Frequency converter | [12 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. |  | K45 | Wet rotor pump | [12 empty] | In accordance with Swiss
|  | K67 | Dry heat exchanger | [12 empty] | Standard
|  | K92 | Compressor piston
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Emergency power supply
| Maintenance and inspection |  | Emergency power supply, UPS / battery system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
An emergency power supply should be provided for the safe operation of a BSL-3 laboratory. Additionally, a UPS / battery system should be planned. Users should be consulted in advance about which installations they require to be linked to the emergency power supply and which to the battery system. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Emergency power supply | [12 empty] | more detail needed
Check functioning of emergency power supply, check for damage, and check environmental conditions Check accumulator and charging device for dirt, damage, corrosion and functioning, and check environmental conditions Clean Check electrolyte Measure cell or block voltage Top up or replace electrolyte |  |  | Accumulator | [12 empty] | in work sheets for
|  |  | Charging device | [12 empty] | individual parts
Cables, pipes, rails
|  |  | Electrical protective measures | [12 empty] | In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Power supply
| Maintenance and inspection |  | Power supply
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
The power supply for BSL-3 laboratories should include at least the following components: general power supply; standby power supply; safety power supply (user-specific); where applicable, generator or accumulators |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Generator | [12 empty] | more detail needed
Check generator for functioning, dirt, damage and corrosion Clean Lubricate bearings Check state of slip rings, carbon brushes and brush holders Check switchgears for dirt, damage and corrosion Clean installation including machine room Check functioning of gearbox and drives Check functioning of electrical and mechanical protective devices |  |  | Accumulators | [12 empty] | in work sheets for
|  |  | Low-voltage switchgear | [12 empty] | individual parts
Reactive current compensation system
|  |  | Electrical subdistributors | [12 empty] | In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Emergency lighting
| Maintenance and inspection |  | Lighting / emergency lighting / safety lighting
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Central battery for safety lighting Individual or group batteries for safety lighting Uninterruptible power supply (UPS) Normal use (e.g. management, active components) |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Central battery | [12 empty] | more detail needed
Check work safety installation for completeness, damage and if necessary for functioning (e.g. voltmeter, actuating rods, replacement fuses, earthing components, warning signs) Check functioning of room lighting and safety lighting Check that electrical connections including connections of busbars are mechanically secured and tighten if necessary |  |  | UPS | [12 empty] | in work sheets for
|  |  | Emergency lighting | [12 empty] | individual parts
Safety lighting
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Telephone system
| Maintenance and inspection |  | Telephone system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A telephone system in BSL-3 containment should include the following components: handsets, switchboard and junction boxes |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Handsets | [12 empty] | more detail needed
Check functioning of handsets, check speech intelligibility Check functioning of loudspeakers Check cables for damage and perfect functioning in junction boxes |  |  | Switchboard | [12 empty] | in work sheets for
|  |  | Junction boxes | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Door opening
| Maintenance and inspection |  | Door intercom and door opening systems
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A door intercom and door opening system in BSL-3 containment should include the following components: loudspeaker, scanners, coder and card reader |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Loudspeaker | [12 empty] | more detail needed
See also telephone system, access control system |  |  | Scanner | [12 empty] | in work sheets for
|  |  | Coder | [12 empty] | individual parts
Card reader
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Electroacoustic alarm
| Maintenance and inspection |  | Electroacoustic alarm system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
An electroacoustic alarm system in BSL-3 containment should include the following components: electroacoustic alarm, movement sensors (including those for access control, intruder detection and CCTV) |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Alarm system | [12 empty] | more detail needed
Check alarm switchboard for dirt, damage and corrosion Check UPS for dirt, damage and corrosion Check loudspeakers and connection boxes for damage Check all loudspeakers for acoustic function, lack of distortion and speech intelligibility Check functioning of amplifiers Check functioning of spare amplifier and standby controls Test alarm activation by making announcement |  |  | Movement sensors | [12 empty] | in work sheets for
individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Hazard alert
| Maintenance and inspection |  | Hazard alert and alarm systems
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A hazard alert and alarm system in BSL-3 containment should include the following components: alarm systems for fire, attack, break-in, manned guarding systems, access control systems, and surveillance systems |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks | [15 empty] | more detail needed
See below Fire alarm system Intruder detection system Access control system | [15 empty] | in work sheets for
individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Fire alarm
| Maintenance and inspection |  | Fire alarm system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
New laboratory buildings should be equipped with a fire alarm system throughout. It should include a PA system for all lab areas, so that instructions can be given in the event of an evacuation. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Switch elements | [12 empty] | more detail needed
Check switch and display elements for dirt, damage and corrosion Check functioning of trigger and display elements in the fire alarm system Check message transfer settings in the fire alarm system |  |  | Trigger elements | [12 empty] | in work sheets for
|  |  | Display elements | [12 empty] | individual parts
Smoke and fire detectors
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Intruder alarm
| Maintenance and inspection |  | Intruder detection system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
An intruder detection system in BSL-3 containment should include the following components: surveillance of external doors; surveillance of ground-floor and first-floor windows; surveillance of entrances to lab area and special areas |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Loudspeaker | [12 empty] | more detail needed
Check loudspeaker system Check functioning of amplifier Check functioning of peripherals Check functioning of loudspeaker circuits Check speech intelligibility Check functioning of trigger and display elements in the rooms Check message transfer settings in the intruder detection system Check display module for functioning, display and control of message transfers |  |  | Amplifier | [12 empty] | in work sheets for
|  |  | Movement sensors | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Access control
| Maintenance and inspection |  | Access control system
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Measures should generally be taken to counter hazards caused by third parties, e.g. vandalism, sabotage or theft. Building security systems (intrusion and hold-up detection systems), intercoms, CCTV, and access control systems cover these requirements. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Scanners | [12 empty] | more detail needed
Check functioning of control panel, check for dirt, damage, corrosion and environmental conditions Check functioning of card reader Check functioning of electrical door openers Check card reader for dirt, damage, corrosion and environmental conditions Check electrical door openers for dirt, damage, corrosion and environmental conditions |  |  | Coder | [12 empty] | in work sheets for
|  |  | Card reader | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## CCTV
| Maintenance and inspection |  | CCTV
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
A CCTV system in BSL-3 containment should include the following components: External CCTV; internal CCTV; movement sensors (including those for access control and intruder detection systems) |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Cameras | [12 empty] | more detail needed
Check video cable for damage and perfect functioning in connection boxes with pivot frames Measure parameters and record all connections from the camera boxes and all other peripheral connection points to central engineering room Check movement sensors for each line for dirt, damage and fastening |  |  | Movement sensors | [12 empty] | in work sheets for
|  |  | Video cable | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Class II MSC
| Maintenance and inspection |  | Class II safety cabinet
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
EN 12 469 certified; ISO 14644-1; depending on type, fixed connection to duct network or discharge into laboratory; locking / control through AC system; locking sliding window with UV light; type with HEPA filter |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Body of safety cabinet | [12 empty] | more detail needed
Check fan for dirt, damage and corrosion, Check differential pressure in HEPA filter; replace filter if necessary, observing safety guidelines Check locking of sliding windows under UV light Check and measure air flow through sliding window into safety cabinet |  |  | HEPA filter | [12 empty] | in work sheets for
|  |  | Fan | [12 empty] | individual parts
UV lighting
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Class III MSC
| Maintenance and inspection |  | Class III safety cabinet
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
EN 12 469 certified, ISO 14644-1 Locking / control through AC system should be observed for Class III safety cabinets Possible combination with Class II safety cabinet advantageous |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Body of safety cabinet HEPA filter Fan | [12 empty] | more detail needed
Check protective gloves for damage and seal Check differential pressure in HEPA filter; replace filter if necessary, observing safety guidelines Check connections for supply and exhaust air to the AC system |  |  | HEPA Filter | [12 empty] | in work sheets for
|  |  | Ventilator | [12 empty] | individual parts
Protective gloves
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Pass-through locks
| Maintenance and inspection |  | Pass-through locks
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Attention should be paid to the following for pass-through locks: controls, gastight doors, body (material + seal), any UV tubes, automatic fumigation device |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Door lock | [12 empty] | more detail needed
Check door locking Check connection for fumigation device Check UV tube lock with door opening |  |  | Body of lock | [12 empty] | in work sheets for
|  |  | any UV tube | [12 empty] | individual parts
Controls
|  |  | Gastight doors | [12 empty] | In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Dunk tanks
| Maintenance and inspection |  | Dunk tanks
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Attention should be paid to the following fo dunk tanks: Locking of dunk tank doors; intercom system where applicable; drain for disinfectant |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Door lock | [12 empty] | more detail needed
Check locking of flaps Check intercom system Check disinfectant drain and clean where applicable |  |  | Body of lock | [12 empty] | in work sheets for
individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Autoclave
| Maintenance and inspection |  | Pass-through autoclaves
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Attention should be paid to the following for pass-through autoclaves for BSL-3 applications: Steam or electric steriliser; sterilisation of condensate; gastight installation (bioseal); door lock; filter for steam and exhaust air (type); temperature range up to 134 °C (prions); volumes for sterilisation |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Body of autoclave | [12 empty] | more detail needed
Check sterilisation using bioindicator Check drainage of condensate water Check door locking Check filter for steam and exhaust air (type) Check temperature range to be used |  |  | Steam steriliser | [12 empty] | in work sheets for
|  |  | Door lock | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed.
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Fumigation
| Maintenance and inspection |  | Fumigation devices
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Generally, separate devices coordinated with the particular ventilation systems. Devices are not considered individually. Please refer to the instructions for use and the declaration of conformity supplied with the equipment. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks | [15 empty] | In accordance with
manufacturer's instructions
Reliability
Safety measure
Disposal
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Material airlocks
| Maintenance and inspection |  | Material airlocks for large equipment
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Attention should be paid to the following for material airlocks: Controls / locking, gastight doors, connections, body (material + seal) |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Door lock | [12 empty] | more detail needed
Check door locking Check connection for fumigation device Check intercom system |  |  | Body of lock | [12 empty] | in work sheets for
individual parts
Controls
|  |  | Gastight doors | [12 empty] | In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………

## Building automation
| Maintenance and inspection |  | Building automation
| Company: |  | Location:
| Person responsible: |  | Facility:
| Place / date: |  | Commissioning:
Function |  | Maintenance / inspection plan
Building control technology should have, at least, interfaces to fire alarm and access control systems, measurement and control technology, and AC systems. For BSL-3 facilties, it is also advisable for information to be relayed to central building control, doorman or safety control centre. |  | Code | Individual part | Task | [11 empty] | Part no.
|  |  |  | Check |  |  | Functioning | Clean | Lubricate | Replace | Vent | Empty | Adjust | Special instructions | Official regulations
|  |  |  | State | Corrosion | Tightness | [9 empty] | Note
Checks |  |  | Master computer | [12 empty] | more detail needed
Check control cabinets, control panels and controls for correct and appropriate installation, and check environmental conditions Check for dirt, damage and corrosion Check protective covers for completeness and fastening Clean Check connections for electrical /mechanical functioning; tighten if necessary Check functional elements (e.g. operation and display units) |  |  | Control cabinet | [12 empty] | in work sheets for
|  |  | Manual operating control | [12 empty] | individual parts
In accordance with
Reliability | [15 empty] | manufacturer's instructions
Safety can be ensured only if the supplier's instructions are observed. | [15 empty] | In accordance with Swiss
Standard
Safety measure
Disposal
Disposal must be carried out in a correct and environmentally sound way. Please consult supplier for advice on correct and environmentally sound disposal.
Notes
|  |  | Key: | T = |  | Daily |  |  | V = |  | Quarterly |  |  | 2J= |  | Biennially
|  |  |  | W = |  | Weekly |  |  | H = |  | Semi-annually |  |  | Z = |  | As required
|  |  |  | M = |  | Monthly |  |  | J = |  | Annually |  |  | A = |  | ……………